Supplementary Methods

1. IHDI grading from predicted landmarks (vector formulation)

Coordinate system. Pixel coordinates are used with origin at the top-left, x to the right, and y downward. For each image we denote:

Right/Left inferior iliac points (Hilgenreiner anchors): p^RI, p^LI

Right/Left lateral acetabular edges: p^RAc, p^LAc

Right/Left proximal femoral metaphyseal borders (lateral, medial): p^RF1, p^RF2, p^LF1, p^LF2

The h-points are defined as the midpoints of the proximal metaphysis:

hR = 0.5 * pRF1 + pRF2

hL=0.5*pLF1+pLF2

Step 1 — Hilgenreiner line 𝓗. The line through p^RI and p^LI is represented as ax + by + c = 0, with:

a = y_LI − y_RI

b = x_RI − x_LI

c = y_RI * x_LI − x_RI * y_LI

Let t = p^LI − p^RI and n = (a, b) be tangent and normal vectors. We normalize n to a unit vector n̂ and orient it cranially (toward decreasing y): if n̂·(0,−1) < 0 then set n̂ ← −n̂.

Step 2 — Perkins base points and lines. Orthogonally project each acetabular edge to 𝓗:

foot(q) = q − ((q − p^RI)·n̂) n̂

p^RP = foot(p^RAc),   p^LP = foot(p^LAc)

Perkins lines 𝓟^R, 𝓟^L are the lines through p^RP, p^LP with direction n̂ (i.e., perpendicular to 𝓗).

Step 3 — Signed positions. For each side S ∈ {R, L}:

Relative to 𝓗 (superior vs inferior): s_H^S = (h^S − p^RI) · n̂

Let t̂ be the unit tangent of 𝓗 (t normalized). Relative to the ipsilateral Perkins line (medial vs lateral):

Right hip: s_P^R = (h^R − p^RP) · t̂  (lateral if s_P^R < −τ, medial if s_P^R > τ)

Left hip:  s_P^L = (h^L − p^LP) · t̂  (lateral if s_P^L >  τ, medial if s_P^L < −τ)

τ is a small tolerance to avoid boundary instability (e.g., a few pixels).

Step 4 — Diagonals (±45° from 𝓗). Let θ be the orientation of 𝓗. Define ipsilateral diagonals by θ − π/4 (right) and θ + π/4 (left). With unit direction d̂^S and cranially oriented normal n̂_D^S = rot(+90°) d̂^S, compute:

s_D^R = (h^R − p^RP) · n̂_D^R

s_D^L = (h^L − p^LP) · n̂_D^L

Step 5 — IHDI grade assignment. Using τ > 0:

Grade 1 (inferomedial): s_H^S ≤ −τ AND medial side

Grade 2 (inferolateral): s_H^S ≤ −τ AND lateral side

Grade 3 (superolateral, at/below diagonal): s_H^S ≥  τ AND lateral side AND s_D^S ≤ τ

Grade 4 (superolateral, above diagonal): s_H^S ≥  τ AND lateral side AND s_D^S >  τ

Tie/degeneracy handling. If ||t|| < ε (degenerate 𝓗), foot(·) is non-finite, or |s_H^S| ≤ τ or |s_P^S| ≤ τ or |s_D^S| ≤ τ (on a boundary), automated grading is withheld and the case is routed to human review. Recommended ε ≈ 1e−6 and τ a few pixels.

2. Agreement analysis for ordinal grades (quadratic-weighted κ)

Statistic. Agreement between AI predictions and the consensus IHDI grade (primary analysis) was summarized by quadratic-weighted Cohen’s κ using agreement weights:

w_ij = 1 − ((i − j)^2) / ((k − 1)^2), with k = 4

Let C = (c_ij) be the confusion matrix with rows as the reference (consensus) and columns as AI predictions, n = Σ_ij c_ij, p_ij = c_ij / n, and p_i·, p_·j the corresponding marginals. Then:

P_o = Σ_ij w_ij p_ij

P_e = Σ_ij w_ij p_i· p_·j

κ_w = (P_o − P_e) / (1 − P_e)

Confidence intervals. We computed 95% confidence intervals by non-parametric bootstrap with 5,000 resamples at the hip level (sampling with replacement and recomputing κ_w in each replicate) and reported percentile intervals (2.5th–97.5th). Random seed was 42. As a sensitivity analysis, a patient-cluster bootstrap (resampling patients and retaining all hips per patient) may also be performed if required by reviewers.

Implementation and cross-check. The primary implementation followed the matrix formulation above; results were cross-checked with scikit-learn’s cohen_kappa_score(y_true, y_pred, labels=[1,2,3,4], weights='quadratic'), yielding identical values. (Python 3.10; NumPy / scikit-learn; exact versions recorded in the accompanying scripts.)

3. Operational quality control (QC) and fail-safe policy

A study is automatically flagged when any predicted landmark has a confidence score < 0.50 or when post-processing detects invalid geometry (e.g., coincident acetabular landmarks or non-finite angles). For flagged studies, model outputs (landmark overlays and auxiliary lines) are displayed to clinicians but automated angle/IHDI-based triage decisions are withheld; the case defaults to “review required” for human assessment. If only one hip passes QC, Rule 1 (both hips IHDI-1 → “no referral”) is not applied; the patient remains “review required”.

4. Computational efficiency: timing protocol (brief)

All 150 test images processed sequentially (batch size = 1) with identical preprocessing and post-processing in GPU and CPU runs.

Half-precision and graph compilation were disabled; identical model weights were used in both runs.

For CPU timing, PyTorch threads were limited to two.

Per-image times were summarized as mean and distribution across the 150 images; environment and exact commands are provided with the scripts.

5. Reproducibility and software

All scripts for preprocessing, inference (including IHDI grading), agreement analysis, and plotting are provided with the Supplementary materials.

Random seeds and configuration files are included to reproduce splits and statistics.

Software environment: Python 3.10; NumPy, SciPy, scikit-learn, Matplotlib; exact versions recorded in the scripts.